## Slide 1: A new solution to plasma chamber cooling for high power ECR ion source operation
J. W. Guo
L. B. Li, J. D. Ma, L. X. Li, Y. C. Feng, W. Lu, W. H. Zhang, X. Fang, X. Z. Zhang, Y. Y. Yang, Y. J. Zhai, W. Huang, D. Hitz, L. T. Sun and H. W. Zhao
Institute of Modern Physics (IMP), Chinese Academy of Sciences
Lanzhou, 730000, China
The 24th International Workshop on ECR Ion Source
East Lansing, USA, September 28th-30th 2020

## Slide 2: Outline
Introduction
Analysis of microchannel cooling for plasma chamber
On-line high power test
Summary

## Slide 3: Evolution of ECRIS Generations
4th G.
O6+
Ar12+
Xe27+
U34+
2nd G.
3rd G.
6.4~18 GHz
24~28 GHz
40~60 GHz
Caprice
ECR4
AECR-U
SERSE
GTS
LECR3
…
VENUS
SECRAL
SECRAL-II
SuSI
SCECR
…
Frequency
Power 0.5-2 kW 5-12 kW 15-30 kW

## Slide 4: Power Scaling Laws
SECRAL Xe27+ and Xe30+ scale linearly with power
TE01 ,  waveguide 32mm
TE01 ,  waveguide 20 mm

## Slide 5: SECRAL-II 28 GHz Superconducting ECR Ion Source
Parameters | SECRAL-II
ωrf (GHz) | 18-28
Axial Field Peaks (T) | 3.7 (Inj.), 2.2 (Ext.)
Mirror Length (mm) | 420
Br at r=63 mm (T) | 2.06
Magnet Cooling | LHe bathing
Dynamic cooling power (W) | ~6
Warm bore ID (mm) | 142 .0
Chamber ID (mm) | 125.0（Br=1.96T）
High power operation
 | SECRAL-II（eμA）
Ar16+ | 620
Ar17+ | 130
Xe30+ | 365
Xe34+ | 110
Frequency ： 28 + 18 GHz
Max. Power ：10 + 2 kW

## Slide 6: Plasma chamber burnout
The plasma chamber cooling has become a serious obstacle for high power operation
Injection: 2 times
Extraction: 7 times
SECRAL-II (2019) 6.5kW (24GHz)+1.7kW (18GHz)
VENUS chamber was burned out twice
（D. Z. Xie, ECRIS2016）

## Slide 7: Origin of plasma chamber burnout
Ave.
Extraction: -2.0%
6 points where the magnetic field is weakest
SECRAL-II (machining error during wire winding)
P=4 kW，1x1 mm2 pixel
T. Thuillier, et al., Rev. Sci. Instrum. 87, 02A736 (2016)
Minimum B field

## Slide 8: Main reason of plasma chamber burnout Very high heat flux
Power ：12 kW
Power density：≈ 1 kW/cm2 Very high heat flux !!!
Z=20 mm
θR=8 mm
Obvious over heating mark in at the surface of cooling water channel.
>160℃
The overall cooling capacity is sufficient to remove much more than 12 kW power. (Total flow rate:15 L/m, ΔT=11.4 ℃）
Water flow
Size of overheating area ＜ 8 * 20 mm
Water pressure：6 bar
Boiling point：≈150 ℃

## Slide 9: Origin of wall destruction Transition from nucleate boiling to film boiling
CHF (Critical heat flux) point at Twall ≈ 30 + 150 = 180 ℃
Nukiyama curve for Pool boiling
Nucleate boiling
film boiling
Bubbles coalescence surface dryout 
Rapid surface temperature increase

## Slide 10: Thermal testing platform
oxy-acetylene flame

## Slide 11: The burnout process of plasma chamber Transition from nucleate boiling to film boiling
As the heat flux is too high, the local water boils completely.
The heat exchange decreases and the aluminum temperature rises rapidly，then burnout.

## Slide 12: Cooling & Structure Optimization Goal : to enhance heat transfer
Plasma flux to the water channel
Plasma flux to the solid aluminum (rotating 30°of the chamber)
Plasma flux to the Additional bar
Water flow
Plasma flux
Additional aluminum bar
Thuillier Proposed an hyper-vapotron cooling schemes ，but it requires a total thickness for the plasma chamber of ~20 mm.
T. Thuillier, et al., Rev. Sci. Instrum. 87, 02A736 (2016)
a
b
c
Notes: Need more efficient cooling solutions to deal with high heat flux！！！

## Slide 13: Microchannel cooling
Features:
Very thin thickness
High power flux (0.5-2 kW/cm2)
Acceptable pressure drop
Micro Channel Cooling seems to be an excellent solution
Technically it means using a micro-fluidics device as heat exchanger for the thermal management of a heat source
Microchannel：D< 1 mm
D.B. Tuckerman, R.F.W. Pease, High performance heat sink for VLSI,IEEE Electron Dev.Lett.EDL-2(5)(1981)126-129.
Power
Heat exchange
coefficient ↑
Surface ↑
Temperature difference
Between water and metal

## Slide 14: Plasma chamber with microchannel cooling
Water flow
52 mm
The microchannel structure is set up in the 6 high heat flux poles
Water channel thickness
Chamber wall thickness
Pressure loss and flow velocity
Water flow rate
Parameters that need to be optimized
15.6 * 15 mm
Fins: 0.4 mm*19
Channel: 0.4 mm *20

## Slide 15: Simulation conditions
Uniform power density distribution
Power density：1 kW/cm2，Surface: 100 mm²（φ11.28 mm），Power : 1 KW
Total power at each pole:≈ 1 kW Max. power density：≈ 1 kW/cm2
Gaussian power density distribution
Max. power density：1 kW/cm2 ，Surface: 314 mm²（φ20 mm）, Power : 1 KW
 | Surface (cm2) | Power (W) | Power density (MW/m2)
φ1 | 0.79 | 8.14 | 10.3038
Φ2-φ1 | 2.36 | 24.04 | 10.1864
Φ4-φ2 | 9.42 | 90.72 | 9.63057
Φ6-φ4 | 15.71 | 133.52 | 8.49905
Φ8-φ6 | 21.99 | 155.12 | 7.05412
Φ10-φ8 | 28.27 | 155.53 | 5.50159
Φ12.5-φ10 | 44.18 | 170.29 | 3.85446
Φ15-φ12.5 | 54 | 128.24 | 2.37481
Φ20-φ15 | 137.45 | 134.39 | 0.97774
total | 314 | 1000 | 

## Slide 16: Water channel thickness
Chart: Water channel thickness (mm)
Chamber | 422 | 255 | 192 | 186 | 179 | 174 | 165 | 156
Water | 350 | 176 | 109 | 103 | 95 | 88 | 78 | 69
Al 6061-T6, Flow rate: 7 L/min, Chamber wall thickness: 1.5 mm, Channel height：1 mm
Uniform power density distribution

## Slide 17: Temperature distribution
Al 6061-T6, Flow rate: 7 L/min, Chamber wall thickness: 1.5mm
1
2
1
2
1
2
 | Power density distribution | T max chamber | T max water
1 | Uniform (φ11.28 mm) - 1 kW/cm2 | 174℃ | 88℃
2 | Gaussian (φ20mm)- 1 kW/cm2 | 159℃ | 78℃
Cross section
Inside-wall
Water
Microchannel: 0.4 mm *20, Channel height：1 mm
Plasma
Plasma

## Slide 18: Chamber wall thickness
Chart: Chamber wall thickness (mm)
Chamber | 195 | 174 | 149
Water | 83 | 88 | 92
1.5
2.0
Al 6061-T6, Flow rate: 7 L/min,
Microchannel: 0.4 mm *20, Channel height：1 mm
Uniform power density distribution
Plasma
Plasma
Plasma

## Slide 19: Water pressure drop & flow velocity
Microchannel: 0.4 mm *20, Channel height：1 mm, Flow rate: 7 L/min
Maximum flow velocity: 18 m/s
Pressure drop: 0.11 Mpa
Usual channel
Maximum flow velocity: 2.6 m/s
Pressure drop: 0.009 Mpa
(52*1*150mm)

## Slide 20: Water flow rate
Uniform power density distribution
Chart: Water flow rate of each channel ( L/ m)
| 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
Chamber | 223 | 204 | 193 | 184 | 178 | 174 | 170 | 166
Water | 141 | 121 | 109 | 100 | 93 | 88 | 83 | 80
Microchannel: 0.4 mm *20, Channel height：1 mm, Chamber wall thickness: 1.5 mm

## Slide 21: Strength analysis
Maximum pressure：2.0 MPa
Al 6061-T6 Yield strength: 100 MPa (@200 ℃);
Maximum yield strength:104 Mpa
Maximum deformation: 0.12 mm

## Slide 22: Plasma chamber with microchannel cooling
Water pressure (Bar) | Flow rate Channel A (L/m) | Flow rate Channel B (L/m) | Flow rate Channel C (L/m)
8.9 | 4.3 | 4.2 | 4.6

## Slide 23: High power operation
Xe34+=106 eμA Io= 2.7 emA
HV= 20 kV, Slit-x= 20 mm
B: 3.41 T, 0.6 T, 1.8 T, 1.87 T
2020/09/18
SECRAL-II operate at the maximum power of the microwave system
ωrf: 24 +18 GHz
Pw= 9.3 + 1.7 kW

## Slide 24: High power operation
Reliable operation at total power 11 kW for more than 48 hours, including 24 hours for Xe34+

## Slide 25: Summary
The efficiency of microchannel cooling embedded in plasma chamber of ECRIS at high power operation is demonstrated at IMP
The breakthrough of plasma chamber cooling can make high power (>10 kW) operation reliable
Microchannel cooling technology is the promising solution for chamber cooling of the fourth-generation ECR ion source with power of 20 kW

## Slide 26
Thanks for your attention!